## Coord System
Improved Off-Axis Winston-Lutz Calculator | [8 empty] | Version 1.4
Matrix Version
| Methodology described here: |  |  |  | Eagle, A, Tallhamer, M, Keener, J, Geneser, S. A simplified and effective off-axis Winston–Lutz for single-isocenter multi-target SRS. J Appl Clin Med Phys. 2022;e13816. https://doi.org/10.1002/acm2.13816
Step 1. Specify coordinate system axes.
Specify the axes for your TPS coordinate system. Use the provided diagram, and identify which TPS axis corresponds to A, B, and C.
|  | A Axis: | +X | [6 empty] | +X | X | TRUE | TRUE | 1
|  | B Axis: | +Y | [6 empty] | -X | Y | TRUE |  | 1
|  | C Axis | +Z | [6 empty] | +Y | Z | TRUE |  | 1
-Y
|  | Duplicate axes selected. Please ensure that each axis is only selected once. | [7 empty] | +Z
-Z
Step 2. Specify sign of rotations.
Specify whether angular rotations are positive or negative for your Treatment Planning System.
This can be determined empirically by using the right-hand rule (RHR).
Note that rotations are positive when clockwise while viewing the axis from negative to positve (i.e. pointing away from you). For this determination, use the A, B, and C axes as shown above.
|  | Couch Angle: |  | ( Coord. System = BEV ) |  | Increasing couch values represent a… | Positive | rotation. |  | 1 | 1 |  |  |  | <select>
|  | Gantry Angle: |  | ( Coord. System = TPS ) |  | Increasing gantry values represent a… | Negative | rotation. |  | -1 | -1 |  |  |  | Positive | 1
|  | Collimator Angle: |  | ( Coord. System = BEV ) |  | Increasing coll. values represent a… | Positive | rotation. |  | 1 | 1 |  |  |  | Negative | -1
Go to "Calcs" tab…

## Calcs
Improved Off-Axis Winston-Lutz Calculator | [12 empty] | Version 1.4
Matrix Version
| Methodology described here: |  |  |  | Eagle, A, Tallhamer, M, Keener, J, Geneser, S. A simplified and effective off-axis Winston–Lutz for single-isocenter multi-target SRS. J Appl Clin Med Phys. 2022;e13816. https://doi.org/10.1002/acm2.13816
Step 3. Enter target coordinates (up to 5 targets).
|  | Step 1.a Enter absolute coords of targets |  |  |  |  | Step 1.b Choose which target is at iso, and which to use for OAWL target.
|  |  | Absolute Coords (cm) |  |  |  | Isocenter Target: |  | 1
|  | Target | x | y | z |  | Off-Axis Target: |  | 5 | [5 empty] | A | B | C
|  | 1 | -2.5 | 3 | -1 | [8 empty] | 1 | -2.5 | 3 | -1 |  | Iso | A | B | C
|  | 2 | -1.5 | 1.5 | 1 | [8 empty] | 2 | -1.5 | 1.5 | 1 |  |  | -2.5 | 3 | -1
|  | 3 | 0 | 0 | 0 |  | Relative Coords (cm) | [6 empty] | 3 | 0 | 0 | 0
|  | 4 | 2 | -1 | -2 |  | X | Y | Z | Dist |  |  |  | 4 | 2 | -1 | -2 |  | Target | A | B | C
|  | 5 | 3 | -2 | 2 |  | 5.5 | -5 | 3 | 8.0156 |  |  |  | 5 | 3 | -2 | 2 |  |  | 3 | -2 | 2
Step 4. Enter desired Couch, Gantry, and initial Collimator angles
Note that the Collimator angle will ultimately be slightly different than desired, as we use a collimator angle correction to align MLC leaf-interfaces.
|  | Couch Angle |  | 30 |  | Gantry Angle |  | 20 |  | Collimator Angle |  | 10
Step 5. Calculate coordinate transformations, to obtain the exact position of target in beam's eye view.
See Details tab for step by step coordinate transformations.
|  | XBEV,target = | 2.4041 |  | YBEV,target = | 5.3419 | [7 empty] | M120 | 0.5 | 8.5 | 1
[14 empty] | HD120 | 0.25 | 3 | 0.5
| Step 6. Determine closest leaf-interfaces, and calculate needed collimator angle correction to align each leaf-interace with target. | [12 empty] | Y Coord. | Gap Between Leaves… |  |  | Y Coord. | Gap Between Leaves…
Yint is the Y coordinate to the interface between the two MLC leaves.
[14 empty] | -19 | 59 | 60 |  | -10.5 | 59 | 60
|  | MLC Type |  | HD120 | 2 | Leaf Thickness (cm) |  | 0.5 | Note: The calculated field falls beyond small MLC leaves, so larger leaf size is used. |  |  |  |  | -18 | 58 | 59 |  | -10 | 58 | 59
[14 empty] | -17 | 57 | 58 |  | -9.5 | 57 | 58
[14 empty] | -16 | 56 | 57 |  | -9 | 56 | 57
|  | For… | Yint = | 5 |  | New Collimator angle, θc,final = |  |  | 17.1716 |  |  |  |  | -15 | 55 | 56 |  | -8.5 | 55 | 56
[6 empty] | Centered between MLC leaves |  |  | 48 | and | 49 |  |  | -14 | 54 | 55 |  | -8 | 54 | 55
[6 empty] | Centered at X = |  |  | 3.0522 |  |  |  |  | -13 | 53 | 54 |  | -7.5 | 53 | 54
[14 empty] | -12 | 52 | 53 |  | -7 | 52 | 53
[14 empty] | -11 | 51 | 52 |  | -6.5 | 51 | 52
[14 empty] | -10 | 50 | 51 |  | -6 | 50 | 51
|  | or for… | Yint = | 5.5 |  | New Collimator angle, θc,final = |  |  | 5.904 |  |  |  |  | -9.5 | 49 | 50 |  | -5.5 | 49 | 50
[6 empty] | Centered between MLC leaves |  |  | 49 | and | 50 |  |  | -9 | 48 | 49 |  | -5 | 48 | 49
[6 empty] | Centered at X = |  |  | 2.0164 |  |  |  |  | -8.5 | 47 | 48 |  | -4.5 | 47 | 48
[14 empty] | -8 | 46 | 47 |  | -4 | 46 | 47
[14 empty] | -7.5 | 45 | 46 |  | -3.75 | 45 | 46
[14 empty] | -7 | 44 | 45 |  | -3.5 | 44 | 45
[14 empty] | -6.5 | 43 | 44 |  | -3.25 | 43 | 44
[14 empty] | -6 | 42 | 43 |  | -3 | 42 | 43
[14 empty] | -5.5 | 41 | 42 |  | -2.75 | 41 | 42
[14 empty] | -5 | 40 | 41 |  | -2.5 | 40 | 41
[14 empty] | -4.5 | 39 | 40 |  | -2.25 | 39 | 40
[14 empty] | -4 | 38 | 39 |  | -2 | 38 | 39
[14 empty] | -3.5 | 37 | 38 |  | -1.75 | 37 | 38
[14 empty] | -3 | 36 | 37 |  | -1.5 | 36 | 37
[14 empty] | -2.5 | 35 | 36 |  | -1.25 | 35 | 36
[14 empty] | -2 | 34 | 35 |  | -1 | 34 | 35
[14 empty] | -1.5 | 33 | 34 |  | -0.75 | 33 | 34
[14 empty] | -1 | 32 | 33 |  | -0.5 | 32 | 33
[14 empty] | -0.5 | 31 | 32 |  | -0.25 | 31 | 32
[14 empty] | 0 | 30 | 31 |  | 0 | 30 | 31
[14 empty] | 0.5 | 29 | 30 |  | 0.25 | 29 | 30
[14 empty] | 1 | 28 | 29 |  | 0.5 | 28 | 29
[14 empty] | 1.5 | 27 | 28 |  | 0.75 | 27 | 28
[14 empty] | 2 | 26 | 27 |  | 1 | 26 | 27
[14 empty] | 2.5 | 25 | 26 |  | 1.25 | 25 | 26
[14 empty] | 3 | 24 | 25 |  | 1.5 | 24 | 25
[14 empty] | 3.5 | 23 | 24 |  | 1.75 | 23 | 24
[14 empty] | 4 | 22 | 23 |  | 2 | 22 | 23
[14 empty] | 4.5 | 21 | 22 |  | 2.25 | 21 | 22
[14 empty] | 5 | 20 | 21 |  | 2.5 | 20 | 21
[14 empty] | 5.5 | 19 | 20 |  | 2.75 | 19 | 20
[14 empty] | 6 | 18 | 19 |  | 3 | 18 | 19
[14 empty] | 6.5 | 17 | 18 |  | 3.25 | 17 | 18
[14 empty] | 7 | 16 | 17 |  | 3.5 | 16 | 17
[14 empty] | 7.5 | 15 | 16 |  | 3.75 | 15 | 16
[14 empty] | 8 | 14 | 15 |  | 4 | 14 | 15
[14 empty] | 8.5 | 13 | 14 |  | 4.5 | 13 | 14
[14 empty] | 9 | 12 | 13 |  | 5 | 12 | 13
[14 empty] | 9.5 | 11 | 12 |  | 5.5 | 11 | 12
[14 empty] | 10 | 10 | 11 |  | 6 | 10 | 11
[14 empty] | 11 | 9 | 10 |  | 6.5 | 9 | 10
[14 empty] | 12 | 8 | 9 |  | 7 | 8 | 9
[14 empty] | 13 | 7 | 8 |  | 7.5 | 7 | 8
[14 empty] | 14 | 6 | 7 |  | 8 | 6 | 7
[14 empty] | 15 | 5 | 6 |  | 8.5 | 5 | 6
[14 empty] | 16 | 4 | 5 |  | 9 | 4 | 5
[14 empty] | 17 | 3 | 4 |  | 9.5 | 3 | 4
[14 empty] | 18 | 2 | 3 |  | 10 | 2 | 3
[14 empty] | 19 | 1 | 2 |  | 10.5 | 1 | 2

## Details
Improved Off-Axis Winston-Lutz Calculator | [14 empty] | Version 1.4
Matrix Version
| Methodology described here: |  |  |  |  | Eagle, A, Tallhamer, M, Keener, J, Geneser, S. A simplified and effective off-axis Winston–Lutz for single-isocenter multi-target SRS. J Appl Clin Med Phys. 2022;e13816. https://doi.org/10.1002/acm2.13816
And t , g , and c be the table, gantry, and initial collimator angle, respectively.
|  |  | Table (t) = | 30 |  |  |  | Gantry (g) = | 20 |  |  |  | Collimator (c) = | 10
Apply sign change for rotation directions
|  |  | Table (t) = | 30 |  |  |  | Gantry (g) = | -20 |  |  |  | Collimator (c) = | 10
Apply sign change to account for object not moving with rotation
|  |  | Table (t) = | 30 |  |  |  | Gantry (g) = | 20 |  |  |  | Collimator (c) = | -10
General Transform Matrices used…
|  | Affine Orthog Projection Matix | [5 empty] | Affine RB_Couch (couch rotation about A axis) | [6 empty] | Affine RC_Gantry (gantry rotation about C axis) | [5 empty] | Affine RB_Coll (collimator rotation about A axis)
|  | 1 | 0 | 0 | 0 |  |  | 0.86603 | 0 | -0.5 | 0 |  |  |  | 0.93969 | 0.34202 | 0 | 0 |  |  | 0.98481 | 0 | 0.17365 | 0
|  | 0 | 1 | 0 | 0 |  |  | 0 | 1 | 0 | 0 |  |  |  | -0.34202 | 0.93969 | 0 | 0 |  |  | 0 | 1 | 0 | 0
|  | 0 | 0 | 1 | 0 |  |  | 0.5 | 0 | 0.86603 | 0 |  |  |  | 0 | 0 | 1 | 0 |  |  | -0.17365 | 0 | 0.98481 | 0
|  | 0 | 0 | 0 | 1 |  |  | 0 | 0 | 0 | 1 |  |  |  | 0 | 0 | 0 | 1 |  |  | 0 | 0 | 0 | 1
Coordinate Matrices…
|  | Iso Coordinates |  |  |  |  | I Matrix | [5 empty] | Target coordinates |  |  |  | T Matrix |  |  | Target Coord Matrix
|  | A | -2.5 |  |  |  | 1 | 0 | 0 | 2.5 |  |  | A | 3 |  |  | 3 |  |  | 5.5
|  | B | 3 |  |  |  | 0 | 1 | 0 | -3 |  |  | B | -2 |  |  | -2 |  |  | -5
|  | C | -1 |  |  |  | 0 | 0 | 1 | 1 |  |  | C | 2 |  |  | 2 |  |  | 3
[7 empty] | 0 | 0 | 0 | 1 | [6 empty] | 1 |  |  | 1
Perform Affine Rotation Coordinate Transformations…
|  | Couch Rot. ( AffineRB_Couch ● Target Matrix ) | [9 empty] | Gantry Rot. ( AffineRC_Gantry ● Couch Rot. ) | [6 empty] | Coll. Rot. ( AffineRB_Coll ● Gantry Rot. )
| A | 3.2631 | [8 empty] | A | 1.3562 | [5 empty] | A | 2.2643
| B | -5 | [8 empty] | B | -5.8145 | [5 empty] | B | -5.8145
| C | 5.3481 | [8 empty] | C | 5.3481 | [5 empty] | C | 5.0313
|  | 1 | [9 empty] | 1 | [6 empty] | 1
Perform Coordinate Transforms for projection magnification…
|  | Final B coordinate |  |  |  | Scaling Factor… |  |  |  |  | Affine Scaling Matrix | [5 empty] | BEV Coordinates ( Affine_Scale ● { Ortho_Proj_M ● m_targ_coll } )
|  | -5.8145 |  |  |  | 1.0617 |  |  |  |  | 1.0617 | 0 | 0 | 0 |  |  | A | 2.4041 | [6 empty] | 1 | 0 | 0 | 0
[11 empty] | 0 | 1.0617 | 0 | 0 |  |  | B | -6.1735 | [6 empty] | 0 | 1 | 0 | 0
[11 empty] | 0 | 0 | 1.0617 | 0 |  |  | C | 5.3419 | [6 empty] | 0 | 0 | 1 | 0
[11 empty] | 0 | 0 | 0 | 1 |  |  | h | 1 | [6 empty] | 0 | 0 | 0 | 1
Final BEV Coordinates
|  |  | XBEV,target = | 2.4041 |  | ( BEV Matrix A coordinate ) | [6 empty] | YBEV,target = | 5.3419 | ( BEV Matrix C coordinate )
New Collimator angle…
[17 empty] | For Yint = | 5 |  |  | For Yint = | 5.5
[17 empty] | H = | 5.858 |  |  | H = | 5.858
|  |  | eq. 12) | [13 empty] | YBEV,final = Yint = | 5 |  |  | YBEV,final = Yint = | 5.5
|  |  | eq. 13) | [13 empty] | XBEV,final = | 3.0522 |  |  | XBEV,final = | 2.0164
|  |  | eq. 14) | [13 empty] | qc,corr = | -7.1716 |  |  | qc,corr = | 4.096
|  |  | eq. 15) | [13 empty] | qc,final = | 17.1716 | deg |  | qc,final = | 5.904 | deg

## Version History
Version History
OAWL Calculator created by Anton Eagle, MS. DABR., Advent Health. 2022.
| Version | Change Notes
| 1.0 | Initial version.
| 1.1 | Added initial section for specifying the sign of rotations.
| 1.2 | Added this Version History.
| 1.3 | Reordered eq 12 and 13. Added "sign" to eq. 13 description.
| 1.4 | Renamed "Improved Off-Axis Winston Lutz Calculator". Added ability to specify the coordinate system used. Also, created a new version of this spreadsheet that uses matricies instead of Trig. Functions.
| 1.5 | Renamed B axis to C, and C axis to B.
| 1.6 | Reoriented axes to conform to DICOM orientation, where DICOM X is A, DICOM Y is B, and DICOM Z is C.